Comparing fluorometric methods (in vivo vs extracted) for cyanoHAB monitoring in six Rhode Island ponds

Stephen D. Shivers1, Jeffrey W. Hollister1, Sophie Fournier2, Jakob A. Stankoski1, and Betty J. Kreakie1

1United States Environmental Protection Agency, Atlantic Coastal Environmental Sciences Division, Narragansett, RI, 02882

2Oak Ridge Institute for Science and Education fellow, US Environmental Protection Agency, Atlantic Coastal Environmental Sciences Division, Narragansett, RI, 02882

*Corresponding author: shivers.stephen@epa.gov

Abstract

Harmful algal blooms caused by cyanobacteria (cyanoHABs) are detrimental to human and environmental health and can be difficult to monitor without specialized training and equipment. A variety of instruments have been developed to measure cyanoHAB indicators (i.e., chlorophyll a or phycocyanin) that do not require advanced laboratory processes (e.g., pigment extraction). We compared measurements from five in vivo fluorometers (Turner Trilogy in-vivo module, Turner Fluorosense, Turner Cyanofluor, bbe AlgaeTorch, and bbe Phycoprobe) to results from solvent-based extractions for chlorophyll a and phycocyanin at six different waterbodies in Rhode Island. We found a strong relationship between extracted phycocyanin and in vivo fluorometers (R2 ranging from 0.78-0.96). We found less consistency between in vivo measurements of chlorophyll a and extracted chlorophyll (R2 between 0.34 and 0.82). Phycocyanin results from in vivo fluorometry were also strongly related with cell counts, which implies that phycocyanin measurements from these instruments can be a good proxy for cell counts. Many federal, state, and local entities use cell counts of cyanobacteria to determine when to issue health or contact advisories for waterbodies. Producing accurate cell counts requires highly specialized training/equipment, processing time, and counts can vary greatly between technicians. The results from this study encourage further adoption of phycocyanin fluorescence for cyanoHAB monitoring efforts.

Keywords: cell counts, chlorophyll a, fluorometry, harmful algal bloom, phycocyanin

Introduction

Harmful algal blooms caused by cyanobacteria (cyanoHABs) are likely to increase in a warming world with greater inputs of nutrients by humans (1,2) These blooms may impact both human and environmental health, particularly if toxins are being produced. CyanoHAB blooms are often transient in nature, can be difficult to accurately quantify, and occur in diverse waterbodies across large spatial scales (3). Despite occuring on large scales, effects of blooms have primarily local impacts and can differ in species composition and toxicity based on local conditions (4,5)). As large, centrally managed monitoring efforts are expensive and time consuming to operate, a substantial portion of cyanoHAB monitoring occurs on a smaller, more local level. As neighborhood and lake associations generally do not have the equipment or training to perform extractive analysis of chlorophyll a and/or phycocyanin (pigments found in all algae and specifically cyanobacteria, respectively), alternative measurement equipment has been developed by several companies.

A variety of instruments have been designed to measure chlorophyll and/or phycocyanin using in vivo fluorescence without an extraction step. These instruments distinguish between the pigments based on their differing emission spectra. Some instruments (i.e., PhycoProbe) use multiple LEDs at varying wavelengths to cause excitation and emission of different algal type pigments, thus creating a finer scale breakdown of community composition (6). These instruments can be used in the field, the lab, a combination of both, and have a wide range in cost. Field instruments may be placed directly in a waterbody or require sample water to be poured into a cuvette and then placed into a handheld fluorometer. In vivo fluorescence can be affected by nonphotochemical quenching causing an underestimation of RFUs as photosystems become saturated with light (7). Estimation problems can also result from differing phytoplankton/cyanobacteria community composition and growth stage, particularly if colonial cyanobacteria are present as the interior of large colonies will not fluoresce as readily as the exterior (8). Additionally, the physical properties of the sample water can greatly affect fluorescence readings based on turbidity levels and the presence of colored dissolved organic matter (CDOM). Some instruments offer a yellow substance correction to help account for CDOM while others do not (9)).

Overall, the physical properties and phytoplankton community composition of a waterbody can cause instruments to estimate pigment concentrations differently. How the different instruments handle these various potential situations can affect their accuracy and precision, and the ability to compare across instrument types. The goal of this paper is to compare the performance of several handheld and in situ fluorometers as tools for cyanoHAB monitoring with respect to a common benchmark (extracted chlorophyll and phycocyanin) across a variety of waterbodies.

Methods

Equipment used for comparison

For this study, we used 5 fluorometers that are commonly used by our collaborators in the northeast US (Table 1). In decreasing order of cost, they are: bbe PhycoProbe, bbe AlgaeTorch, Turner Trilogy (in vivo chlorophyll and orange modules), Turner Cyanofluor, and Turner FluoroSense. Each instrument was calibrated according to manufacturer specifications to allow a comparison based on how each instrument is intended to be used “out of the box”.  Measurements from these fluorometers were compared with the results of a solvent-based extraction for chlorophyll and phycocyanin using a Turner Trilogy with a chla-na and orange module, respectively.

Table 1:  Manufacturer specifications for the fluorometers that were used in this study.

| Fluorometer | Analysis Method | Measurement Range | Resolution/MDL |
| --- | --- | --- | --- |
| bbe PhycoProbe | cuvette | 0-200 µg chl-a/L | 0.01 µg chl-a/L |
| bbe AlgaeTorch | in situ | 0-500 µg chl-a/L | 0.15 µg chl-a/L |
| Turner Trilogy (in vivo) | cuvette | 0-300 µg/L | 0.025 µg/L |
| Turner Cyanofluor | cuvette | 0-100 µg/L | 0.3 µg/L |
| Turner FluoroSense | in situ | 0-199 µg/L | 1 µg/L |

Field Sampling Methods

Field samples were collected from 6 ponds in Rhode Island (Figure 1). Samples from Barber, Indian, and Yawgoo were collected on 2021-10-06 and samples from Curran, Mashapaug, and Warwick were collected on 2021-10-22. Surface samples were collected in triplicate by wading into each pond to a minimum depth of 2 feet to avoid collecting near the sediment. If a surface scum was present, the scum was gently brushed aside to avoid being collected in the sample bottles. Samples were collected in acid-washed 1 L amber bottles, placed in a cooler with ice, and stored at 4° C until analysis. At the same location as sample collection, measurements of chlorophyll and phycocyanin were made using the AlgaeTorch and the FluoroSense.

Fig 1: Six field sampling locations in Rhode Island, USA

Analysis of field samples

Within 24 hours of collection, field samples were analyzed using the PhycoProbe, CyanoFluor, and the Trilogy (in-vivo), as well as filtered onto 0.7 µm pre-ashed glass fiber filters and frozen at -20° C for solvent-based extraction. Sample water filtered through a 0.22 µm syringe filter was used to correct for yellow substances on the PhycoProbe, CyanoFluor, and Trilogy in vivo module. Chlorophyll a was extracted by placing frozen GF/F filters in centrifuge tubes with 90% acetone, then placing the tubes for a 20-minute period in a sonicating water bath to lyse cyanobacterial cells (10). Phycocyanin was extracted from frozen GF/F filters using a 50 mM phosphate buffer as the extraction solvent and following a 15-minute period in a sonicating water bath to lyse cyanobacterial cells (11). Extracted samples were analyzed using a Turner Trilogy with a chla na module (chlorophyll) and an orange module (phycocyanin).

Phytoplankton Counts

All phytoplankton samples were collected on 2021-10-06. Water samples for phytoplankton identification and enumeration were decanted from field collected samples into 250 ml amber HDPE bottles and were preserved with 25% glutaraldehyde (0.25% working concentration). Samples were stored at 4° C and shipped the week after collection on ice to Phycotech, Inc. A minimum of 400 natural units were counted per sample and reported as the cell count. A full description of methods and a quality assurance plan are available from Phycotech, Inc. (http://www.phycotech.com).

Data Analysis

To compare extracted concentrations and cell counts to the various instruments we used simple linear regression and scatterplots. We derived the coefficient of determination from the regressions as a measure of fit and also qualitatively compared the device measurements to the extracted concentrations and phytoplankton counts with scatterplots. We do not compare slopes nor look for a slope of 1 as we have little expectation that the extracted concentrations would equal the values from the various devices given the difference in units and the different RFU returns across the devices. All data and code for this analysis are available via GitHub at https://github.com/usepa/fluoroproj.  This repo has also been archived at Zenodo (12).

Results

Chlorophyll

Extracted chlorophyll measurements ranged from a mean of 1.3 µg/L in Indian Lake to a mean of 32.2 µg/L in Mashapaug Pond (Table 2). The Algaetorch and Phycoprobe measurements exhibited a strong relationship with extracted chlorophyll concentrations with R2 values of 0.82 and 0.81, respectively (Figure 2).  The Cyanofluor (R2=0.35) and Trilogy in vivo (R2=0.34) measurements increased as extracted chlorophyll concentrations increased, but the relationship was not as strong compared to the other instruments.

Table 2: Summary (mean and standard deviation) of extracted chlorophyll and phycocyanin from Turner Trilogy for waterbodies sampled in Rhode Island.

| waterbody | chlorophyll (µg/L) | phycocyanin (µg/L) |
| --- | --- | --- |
| Barber | 4.5 ± 0.2 | 0.2 ± 0.0 |
| Curran | 19.6 ± 0.2 | 1.5 ± 0.4 |
| Indian | 1.3 ± 0.0 | 0.2 ± 0.0 |
| Mashapaug | 32.2 ± 3.6 | 10.5 ± 2.6 |
| Warwick | 20.0 ± 1.0 | 2.4 ± 0.5 |
| Yawgoo | 2.8 ± 0.2 | 1.0 ± 0.1 |

Fig 2: Comparison of fluorometer measurements of chlorophyll by waterbody. Blue lines represent the best fit regression line. R2 on plot is the coefficient of determination for this line.

Phycocyanin

Extracted phycocyanin measurements ranged from a mean of 0.2 µg/L in Indian Lake and Barber Pond to 10.5 µg/L in Mashapaug Pond (Table 2). The Algaetorch (R2=0.86) and Phycoprobe (R2=0.84) were strongly related to extracted phycocyanin concentrations (Figure 3). Phycocyanin measurements for the Fluorosense (R2=0.96) and the Cyanofluor (R2=0.78) were also strongly related with extracted phycocyanin.

Fig 3: Comparison of fluorometer measurements of phycocyanin by Rhode Island waterbody. Blue lines represent the best fit regression line. R2 on plot is the coefficient of determination for this line.

Phytoplankton Counts

Cyanophytes were the dominant taxa at all waterbodies with the highest concentrations observed at Curran (38,019 natural units/mL), Mashapaug (89,806 natural units/mL), and Warwick (45,847 natural units/mL)(Figure 4).  Total concentrations were highest at Mashapaug (92,527 natural units/mL), Warwick (70,728 natural units/mL), and Curran (40,428 natural units/mL).

Fig 4: Cell counts (natural units/mL) of phytoplankton samples from six waterbodies in Rhode Island.

Extracted chlorophyll a exhibited the strongest relationship with cyanobacterial cell counts (R2=0.92) followed by the Algaetorch (R2=0.78) and the Phycoprobe (R2=0.59). The Cyanofluor (R2=0.12) and Trilogy in vivo module (0.11) exhibited a weaker relationship (Figure 5). Phycocyanin measurements from all instruments were strongly related to cell counts (R2 between 0.9 and 0.97)(Figure 6).

Fig 5: Comparison of fluorometer chlorophyll to cyanobacteria cell counts (natural units/mL). Blue lines represent the best fit regression line. R2 on plot is the coefficient of determination for this line.

Fig 6: Comparison of fluorometer phycocyanin to cyanobacteria cell counts (natural units/mL). Blue lines represent the best fit regression line. R2 on plot is the coefficient of determination for this line.

Discussion

All fluorometers performed well when comparing phycocyanin fluorescence measurements to extracted measurements and these results agree with previous studies that supported the use of phycocyanin as an indicator of HABs caused by cyanobacteria (13–15). Phycocyanin fluorescence from fluorometers was also strongly related to cyanobacterial cell counts. Cell counts of cyanobacteria, and not total cell counts, were used because cyanobacterial cell counts are commonly used to define a bloom and as HAB advisory criteria (16,17).  There were differences between fluorescence concentrations and extracted concentrations that may have been the result of either overestimation by the fluorometers or underestimation because of an inefficient extraction process. Neverthless, any of these instruments could be used to effectively monitor for the presence of cyanobacteria by measuring more direct parameters of cyanobacteria than chlorophyll a.

The Algaetorch and Phycoprobe had strong relationships with extracted chlorophyll and cyanobacterial cell counts likely because those instruments excite at multiple wavelengths which better estimates chlorophyll content of cyanobacteria (18,19). Weaker relationships by the Cyanofluor and Trilogy in vivo module were driven by the underestimation of chlorophyll in Mashapaug Pond. Cell counts of cyanobacteria were the highest in Mashapaug and it is likely that these high concentrations were causing underestimation of the in vivo concentrations of chlorophyll due to intracellular chlorophyll storage (19). Extracted chlorophyll had the strongest relationship with cyanobacteria cell counts which follows as extracted chlorophyll has been used as a proxy for phytoplankton biomass for many years.

This study supports the use of commercially available fluorometers for monitoring cyanoHABs when other techniques are not feasible.  In lieu of extractive techniques, in vivo chlorophyll is typically used for routine monitoring of cyanoHABs, but this study found in vivo phycocyanin to be better related to cyanobacterial cell counts across a variety of instruments and is a more direct measure of cyanobacteria than chlorophyll. Therefore, measuring in vivo phycocyanin may be a better solution for HAB monitoring as several other studies have also demonstrated (20–22). There are a variety of instruments capable of measuring phycocyanin using in vivo fluorometry across a spectrum of prices. We hope this study will encourage further adoption of using in vivo phycocyanin as alternative to cell counts and in situations where extractive techniques are not able to be used.

Acknowledgments

We are grateful to Hilary Snook, Darryl Keith, Steve Rego, Kara Godineaux, Marty Chintala, and Tim Gleason for constructive reviews of this paper. The views expressed in this article are those of the authors and do not necessarily represent the views or policies of the U.S. Environmental Protection Agency. Any mention of trade names, products, or services does not imply an endorsement by the U.S. Government or the U.S. Environmental Protection Agency. The EPA does not endorse any commercial products, services, or enterprises. This contribution is identified by the tracking number ORD-060125 of the Atlantic Coastal Environmental Sciences Division, Office of Research and Development, Center for Environmental Measurement and Modeling, US Environmental Protection Agency. Lastly, special thanks to Gracie, Anastasia, and Lilac for cheer leading and support during challenging analysis times.

References

1.	Paerl HW, Scott JT, McCarthy MJ, Newell SE, Gardner WS, Havens KE, et al. It Takes Two to Tango: When and Where Dual Nutrient (N & P) Reductions Are Needed to Protect Lakes and Downstream Ecosystems. Environmental Science & Technology. 2016 Oct;50(20):10805–13. doi:10.1021/acs.est.6b02575

2.	Schindler DW, Hecky RE, Findlay DL, Stainton MP, Parker BR, Paterson MJ, et al. Eutrophication of lakes cannot be controlled by reducing nitrogen input: Results of a 37-year whole-ecosystem experiment. Proceedings of the National Academy of Sciences. 2008 Aug;105(32):11254–8. doi:10.1073/pnas.0805108105

3.	Wynne TT, Stumpf RP, Tomlinson MC, Dyble J. Characterizing a cyanobacterial bloom in Western Lake Erie using satellite imagery and meteorological data. Limnology and Oceanography. 2010 Sep;55(5):2025–36. doi:10.4319/lo.2010.55.5.2025

4.	Chaffin JD, Davis TW, Smith DJ, Baer MM, Dick GJ. Interactions between nitrogen form, loading rate, and light intensity on Microcystis and Planktothrix growth and microcystin production. Harmful Algae. 2018 Mar;73:84–97. doi:10.1016/j.hal.2018.02.001

5.	Gobler CJ, Burkholder JM, Davis TW, Harke MJ, Johengen T, Stow CA, et al. The dual role of nitrogen supply in controlling the growth and toxicity of cyanobacterial blooms. Harmful Algae. 2016 Apr;54:87–97. doi:10.1016/j.hal.2016.01.010

6.	Beutler M, Wiltshire KH, Meyer B, Moldaenke C, Lüring C, Meyerhöfer M, et al. A fluorometric method for the differentiation of algal populations in vivo and in situ. Photosynthesis Research. 2002 Apr;72(1):39–53. doi:10.1023/A:1016026607048

7.	Liu X, Georgakakos AP. Chlorophyll a estimation in lakes using multi-parameter sonde data. Water Research. 2021 Oct;205:117661. doi:10.1016/j.watres.2021.117661

8.	Rousso BZ, Bertone E, Stewart R, Aguiar A, Chuang A, Hamilton DP, et al. Chlorophyll and phycocyanin in-situ fluorescence in mixed cyanobacterial species assemblages: Effects of morphology, cell size and growth phase. Water Research. 2022 Apr;212:118127. doi:10.1016/j.watres.2022.118127

9.	Cremella B, Huot Y, Bonilla S. Interpretation of total phytoplankton and cyanobacteria fluorescence from cross‐calibrated fluorometers, including sensitivity to turbidity and colored dissolved organic matter. Limnology and Oceanography: Methods. 2018 Dec;16(12):881–94. doi:10.1002/lom3.10290

10.	Welschmeyer NA. Fluorometric analysis of chlorophyll a in the presence of chlorophyll b and pheopigments. Limnology & Oceanography. 1994 Dec;39(8):1985–92. doi:10.4319/lo.1994.39.8.1985

11.	Kasinak JME, Holt BM, Chislock MF, Wilson AE. Benchtop fluorometry of phycocyanin as a rapid approach for estimating cyanobacterial biovolume. Journal of Plankton Research. 2015;37(1):248–57.

12.	Hollister J. Comparing fluorometric methods ( in vivo vs extracted) for cyanoHAB monitoring in six Rhode Island ponds [Internet]. 2024 Dec 16. doi:10.5281/ZENODO.14502271

13.	Caraballo YA, Wu M. Summer Phytoplankton Assemblages and Harmful Cyanobacterial Blooms in Freshwaters of New Jersey. Northeastern Naturalist. 2023 Jul 27;30(3). doi:10.1656/045.030.0301

14.	Izydorczyk K, Tarczynska M, Jurczak T, Mrowczynski J, Zalewski M. Measurement of phycocyanin fluorescenceas an online early warning system for cyanobacteria in reservoir intake water. Environ Toxicol. 2005 Aug;20(4):425–30. doi:10.1002/tox.20128

15.	Mchau GJ, Makule E, Machunda R, Gong YY, Kimanya M. Phycocyanin as a proxy for algal blooms in surface waters: case study of Ukerewe Island, Tanzania. Water Practice and Technology. 2019 Mar 1;14(1):229–39. doi:10.2166/wpt.2019.005

16.	Chorus I. Toxic Cyanobacteria in Water: A Guide to Their Public Health Consequences, Monitoring and Management. Welker M, editor. Erscheinungsort nicht ermittelbar: Taylor & Francis; 2021. 1 p.

17.	Graham JL, Loftin KA, Ziegler A, Meyer MT. Guidelines for Design and Sampling for Cyanobacterial Toxin and Taste-and-Odor Studies in Lakes and Reservoirs. USGS; 2008. p. 39. (U.S. Geological Survey Scientific Investigations Report). Report Nos.: 2008–5038.

18.	Gregor J, Maršálek B. Freshwater phytoplankton quantification by chlorophyll a: a comparative study of in vitro, in vivo and in situ methods. Water Research. 2004 Feb;38(3):517–22. doi:10.1016/j.watres.2003.10.033

19.	Zamyadi A, McQuaid N, Prévost M, Dorner S. Monitoring of potentially toxic cyanobacteria using an online multi-probe in drinking water sources. J Environ Monit. 2012;14(2):579–88. doi:10.1039/C1EM10819K

20.	Thomson-Laing G, Puddick J, Wood SA. Predicting cyanobacterial biovolumes from phycocyanin fluorescence using a handheld fluorometer in the field. Harmful Algae. 2020 Jul;97:101869. doi:10.1016/j.hal.2020.101869

21.	Brient L, Lengronne M, Bertrand E, Rolland D, Sipel A, Steinmann D, et al. A phycocyanin probe as a tool for monitoring cyanobacteria in freshwater bodies. J Environ Monit. 2008;10(2):248–55. doi:10.1039/B714238B

22.	McQuaid N, Zamyadi A, Prévost M, Bird DF, Dorner S. Use of in vivo phycocyanin fluorescence to monitor potential microcystin-producing cyanobacterial biovolume in a drinking water source. J Environ Monit. 2011 Feb 9;13(2):455–63. doi:10.1039/C0EM00163E