## Sheet1
Date: 3/31/2021 11:30:15 AM
Device: DFP-01-16
Name | Explanation | Value | Unit | Questions | Notes | Default Value
Common parameters
Warm-up time | During the warm-up time, the LEDs are on, but the measurement has not yet started. | 5 | s | Leave at default? |  | 5
Measurement interval | This is the interval between the start of one measurement to the start of the next measurement. For online measurements without external power supply, the interval can be set to save energy. In this case, please use the power-down option | 0 | s | Wait 10 seconds for ambient light to return to normal? |  | 0
Number of measurements | Performs a given number of measurements and stops afterwards. Only valid if continuous measurement is set to autostop | 3 |  | Is averaging 3 measurements a good idea? Is exitation exhausted after the first? |  | 20
LED measuring time | Duration of the LED brightness measurement. Only valid for PhycoProbes up to series 17 | 20 | 1/10s | What is the significance of this? |  | 20
Measuring type | Select between continuous, single and autostop (after a number of measurements given in “number of measurements" | autostop |  |  |  | continuous
Measurement time | Duration of a single measurement. The longer the time, the more stable the results are. | 60 | 1/10s | Is 6 seconds necessary? |  | 20
Send results | If set, the results of each measurement will be sent to the PC directly after the measurement. Please disable for maximum measuring frequency. In case ‘Send results’ is set to ‘off’, please be aware that you will not receive results directly during measurement to bbe++ even when you started the measurement by way of bbe++. In this case, please download the data from the PhycoProbe later by use of the USB adapter and import the flp file in question to bbe++ | on |  |  |  | on
Power-down after measurement | If set, the PhycoProbe switches off after each measurement to save power. This makes it possible to collect data from some days or weeks depending on the measuring frequency | off |  |  |  | off
Store results | If set, the results will be stored in the internal data logger | on |  |  |  | on
Start measurement on power-up | If set, the PhycoProbe starts a measurement after starting. This is set when using the PhycoProbe as a stand-alone unit without any PC | off |  |  |  | off
Enter comment at start of measurement | If set, a comment can be entered when starting the measurement | on |  |  | Using to enter sample info (pond, site, etc.) | off
Air pressure from first measurement | If set, the air pressure reading for the calculation of the depth is taken from the first dataset. If not set, it has to be measured separately | on |  |  |  | off
Calculate concentrations after each measurement | If set, the PhycoProbe calculates the chlorophyll concentration of each algae class. This can be used if the instrument is operated at a PLC system without Windows. If not set, the concentration is calculated in bbe++. This is faster. | off |  |  |  | on
Serial number | Shows the serial number of the PhycoProbe | 116 |  |  |  | 116
Software version | Shows the software version of the firmware | 274 |  |  |  | 274
Multiple measurements per FLP file |  | on |  |  |  | on
Name of parameter set | Shows the name of the current parameter set. This is only available if there are more than one parameter sets stored in the PhycoProbe | Workstation |  |  |  | Workstation
Fit parameters
Name | Explanation | Value | Unit | Questions | Notes | Default Value
Name of class 1 | The names of the algae classes can be entered here | Green Algae |  |  |  | Green Algae
Name of class 2 |  | Bluegreen |  |  |  | Bluegreen
Name of class 3 |  | Diatoms |  |  |  | Diatoms
Name of class 4 |  | Cryptophyta |  |  |  | Cryptophyta
Name of class 5 |  | Planktothrix |  |  |  | Planktothrix
Name of class 6 |  | #6 |  |  |  | #6
Name of class 7 |  | #7 |  |  |  | #7
Name of class 8 |  | #8 |  |  |  | #8
Name of class 9 |  | #9 |  |  |  | #9
Name of class 10 |  | Turbidity |  |  |  | Turbidity
Name of class 11 |  | unbound Phycocyanin |  |  |  | unbound Phycocyanin
Name of class 12 |  | Yellow substances |  |  |  | Yellow substances
Green Algae enabled for fit | Each algae class can be enabled / disabled. There are 12 possible classes that can be used, but please note that the used classes should be characterised with a unique fingerprint to omit cross talking and therefore wrong differentiation | yes |  |  |  | yes
Bluegreen enabled for fit |  | yes |  |  |  | yes
Diatoms enabled for fit |  | yes |  |  |  | yes
Cryptophyta enabled for fit |  | yes |  |  |  | yes
Planktothrix enabled for fit |  | yes |  |  |  | yes
#6 enabled for fit |  | no |  |  |  | no
#7 enabled for fit |  | no |  |  |  | no
#8 enabled for fit |  | no |  |  |  | no
#9 enabled for fit |  | no |  |  |  | no
Turbidity enabled for fit |  | yes |  |  |  | yes
unbound Phycocyanin enabled for fit |  | yes |  |  |  | yes
Yellow substances enabled for fit | If set to yellow substances (on), the measured concentration is not added to the total algae concentration and the displayed unit is r.u. | yes |  |  |  | yes
Green Algae | The classes will be treated differently depending on their set type. All components marked as algae class will be added to create the total chlorophyll concentration. | Algae class |  |  |  | Algae class
Bluegreen |  | Algae class |  |  |  | Algae class
Diatoms |  | Algae class |  |  |  | Algae class
Cryptophyta |  | Algae class |  |  |  | Algae class
Planktothrix |  | Algae class |  |  |  | Algae class
#6 |  | Algae class |  |  |  | Algae class
#7 |  | Algae class |  |  |  | Algae class
#8 |  | Algae class |  |  |  | Algae class
#9 |  | Algae class |  |  |  | Algae class
Turbidity |  | Turbidity |  |  |  | Turbidity
unbound Phycocyanin |  | Pigment |  |  |  | Pigment
Yellow substances |  | Yellow substances |  |  |  | Yellow substances
Ultrafiltrated water offset of LED 1 [700 nm] det. 1 | Displays the results of the offset calibration of ultra-filtrated sample water | 17.58 | digits |  |  | 17.58
Ultrafiltrated water offset of LED 3 [525 nm] det. 1 |  | 3.0651 | digits |  |  | 3.0651
Ultrafiltrated water offset of LED 4 [570 nm] det. 1 |  | 0.98766 | digits |  |  | 0.98766
Ultrafiltrated water offset of LED 5 [610 nm] det. 1 |  | 0.81351 | digits |  |  | 0.81351
Ultrafiltrated water offset of LED 6 [370 nm] det. 1 |  | 27.617 | digits |  |  | 27.617
Ultrafiltrated water offset of LED 7 [590 nm] det. 1 |  | 1.696 | digits |  |  | 1.696
Ultrafiltrated water offset of LED 8 [470 nm] det. 1 |  | 10.69 | digits |  |  | 10.69
Ultrafiltrated water offset of LED 1 [700 nm] det. 2 |  | 1.2189 | digits |  |  | 1.2189
Ultrafiltrated water offset of LED 3 [525 nm] det. 2 |  | 5.4287 | digits |  |  | 5.4287
Ultrafiltrated water offset of LED 4 [570 nm] det. 2 |  | 0.34607 | digits |  |  | 0.34607
Ultrafiltrated water offset of LED 5 [610 nm] det. 2 |  | 11.409 | digits |  |  | 11.409
Ultrafiltrated water offset of LED 6 [370 nm] det. 2 |  | 29.883 | digits |  |  | 29.883
Ultrafiltrated water offset of LED 7 [590 nm] det. 2 |  | 15.805 | digits |  |  | 15.805
Ultrafiltrated water offset of LED 8 [470 nm] det. 2 |  | 10.754 | digits |  |  | 10.754
Distilled water offset of LED 1 [700 nm] det. 1 | Displays the results of the offset calibration of distilled water. | 18.884 | digits |  |  | 18.884
Distilled water offset of LED 3 [525 nm] det. 1 |  | 0.91982 | digits |  |  | 0.91982
Distilled water offset of LED 4 [570 nm] det. 1 |  | 0.88308 | digits |  |  | 0.88308
Distilled water offset of LED 5 [610 nm] det. 1 |  | 1.0421 | digits |  |  | 1.0421
Distilled water offset of LED 6 [370 nm] det. 1 |  | 4.9165 | digits |  |  | 4.9165
Distilled water offset of LED 7 [590 nm] det. 1 |  | 1.5751 | digits |  |  | 1.5751
Distilled water offset of LED 8 [470 nm] det. 1 |  | 0.90225 | digits |  |  | 0.90225
Distilled water offset of LED 1 [700 nm] det. 2 |  | 1.9457 | digits |  |  | 1.9457
Distilled water offset of LED 3 [525 nm] det. 2 |  | 2.2684 | digits |  |  | 2.2684
Distilled water offset of LED 4 [570 nm] det. 2 |  | 0.26737 | digits |  |  | 0.26737
Distilled water offset of LED 5 [610 nm] det. 2 |  | 12.269 | digits |  |  | 12.269
Distilled water offset of LED 6 [370 nm] det. 2 |  | 4.5961 | digits |  |  | 4.5961
Distilled water offset of LED 7 [590 nm] det. 2 |  | 15.024 | digits |  |  | 15.024
Distilled water offset of LED 8 [470 nm] det. 2 |  | 0.64514 | digits |  |  | 0.64514
Fingerprint of LED 1 [700 nm] det. 1 Green Algae | Displays the results of the algae class calibrations | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 1 [700 nm] det. 1 Bluegreen |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 1 [700 nm] det. 1 Diatoms |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 1 [700 nm] det. 1 Cryptophyta |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 1 [700 nm] det. 1 Planktothrix |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 1 [700 nm] det. 1 #6 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 1 [700 nm] det. 1 #7 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 1 [700 nm] det. 1 #8 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 1 [700 nm] det. 1 #9 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 1 [700 nm] det. 1 Turbidity |  | 3.3558 | digits/Âµg/l |  |  | 3.3558
Fingerprint of LED 1 [700 nm] det. 1 unbound Phycocyanin |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 1 [700 nm] det. 1 Yellow substances |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 3 [525 nm] det. 1 Green Algae |  | 0.32921 | digits/Âµg/l |  |  | 0.32921
Fingerprint of LED 3 [525 nm] det. 1 Bluegreen |  | 0.73641 | digits/Âµg/l |  |  | 0.73641
Fingerprint of LED 3 [525 nm] det. 1 Diatoms |  | 1.8907 | digits/Âµg/l |  |  | 1.8907
Fingerprint of LED 3 [525 nm] det. 1 Cryptophyta |  | 2.4642 | digits/Âµg/l |  |  | 2.4642
Fingerprint of LED 3 [525 nm] det. 1 Planktothrix |  | 4.6197 | digits/Âµg/l |  |  | 4.6197
Fingerprint of LED 3 [525 nm] det. 1 #6 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 3 [525 nm] det. 1 #7 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 3 [525 nm] det. 1 #8 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 3 [525 nm] det. 1 #9 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 3 [525 nm] det. 1 Turbidity |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 3 [525 nm] det. 1 unbound Phycocyanin |  | 0.2801 | digits/Âµg/l |  |  | 0.2801
Fingerprint of LED 3 [525 nm] det. 1 Yellow substances |  | 2.1387 | digits/Âµg/l |  |  | 2.1387
Fingerprint of LED 4 [570 nm] det. 1 Green Algae |  | 0.15638 | digits/Âµg/l |  |  | 0.15638
Fingerprint of LED 4 [570 nm] det. 1 Bluegreen |  | 0.89976 | digits/Âµg/l |  |  | 0.89976
Fingerprint of LED 4 [570 nm] det. 1 Diatoms |  | 0.4064 | digits/Âµg/l |  |  | 0.4064
Fingerprint of LED 4 [570 nm] det. 1 Cryptophyta |  | 1.5051 | digits/Âµg/l |  |  | 1.5051
Fingerprint of LED 4 [570 nm] det. 1 Planktothrix |  | 2.3943 | digits/Âµg/l |  |  | 2.3943
Fingerprint of LED 4 [570 nm] det. 1 #6 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 4 [570 nm] det. 1 #7 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 4 [570 nm] det. 1 #8 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 4 [570 nm] det. 1 #9 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 4 [570 nm] det. 1 Turbidity |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 4 [570 nm] det. 1 unbound Phycocyanin |  | 0.34615 | digits/Âµg/l |  |  | 0.34615
Fingerprint of LED 4 [570 nm] det. 1 Yellow substances |  | 0.12531 | digits/Âµg/l |  |  | 0.12531
Fingerprint of LED 5 [610 nm] det. 1 Green Algae |  | 0.91488 | digits/Âµg/l |  |  | 0.91488
Fingerprint of LED 5 [610 nm] det. 1 Bluegreen |  | 5.3202 | digits/Âµg/l |  |  | 5.3202
Fingerprint of LED 5 [610 nm] det. 1 Diatoms |  | 1.2385 | digits/Âµg/l |  |  | 1.2385
Fingerprint of LED 5 [610 nm] det. 1 Cryptophyta |  | 2.677 | digits/Âµg/l |  |  | 2.677
Fingerprint of LED 5 [610 nm] det. 1 Planktothrix |  | 3.5741 | digits/Âµg/l |  |  | 3.5741
Fingerprint of LED 5 [610 nm] det. 1 #6 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 5 [610 nm] det. 1 #7 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 5 [610 nm] det. 1 #8 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 5 [610 nm] det. 1 #9 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 5 [610 nm] det. 1 Turbidity |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 5 [610 nm] det. 1 unbound Phycocyanin |  | 1.9368 | digits/Âµg/l |  |  | 1.9368
Fingerprint of LED 5 [610 nm] det. 1 Yellow substances |  | -0.2221 | digits/Âµg/l |  |  | -0.2221
Fingerprint of LED 6 [370 nm] det. 1 Green Algae |  | 2.7539 | digits/Âµg/l |  |  | 2.7539
Fingerprint of LED 6 [370 nm] det. 1 Bluegreen |  | 1.5041 | digits/Âµg/l |  |  | 1.5041
Fingerprint of LED 6 [370 nm] det. 1 Diatoms |  | 3.5392 | digits/Âµg/l |  |  | 3.5392
Fingerprint of LED 6 [370 nm] det. 1 Cryptophyta |  | 3.1797 | digits/Âµg/l |  |  | 3.1797
Fingerprint of LED 6 [370 nm] det. 1 Planktothrix |  | 1.6045 | digits/Âµg/l |  |  | 1.6045
Fingerprint of LED 6 [370 nm] det. 1 #6 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 6 [370 nm] det. 1 #7 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 6 [370 nm] det. 1 #8 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 6 [370 nm] det. 1 #9 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 6 [370 nm] det. 1 Turbidity |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 6 [370 nm] det. 1 unbound Phycocyanin |  | 0.37584 | digits/Âµg/l |  |  | 0.37584
Fingerprint of LED 6 [370 nm] det. 1 Yellow substances |  | 22.769 | digits/Âµg/l |  |  | 22.769
Fingerprint of LED 7 [590 nm] det. 1 Green Algae |  | 0.80272 | digits/Âµg/l |  |  | 0.80272
Fingerprint of LED 7 [590 nm] det. 1 Bluegreen |  | 4.6764 | digits/Âµg/l |  |  | 4.6764
Fingerprint of LED 7 [590 nm] det. 1 Diatoms |  | 1.2725 | digits/Âµg/l |  |  | 1.2725
Fingerprint of LED 7 [590 nm] det. 1 Cryptophyta |  | 3.8672 | digits/Âµg/l |  |  | 3.8672
Fingerprint of LED 7 [590 nm] det. 1 Planktothrix |  | 4.258 | digits/Âµg/l |  |  | 4.258
Fingerprint of LED 7 [590 nm] det. 1 #6 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 7 [590 nm] det. 1 #7 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 7 [590 nm] det. 1 #8 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 7 [590 nm] det. 1 #9 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 7 [590 nm] det. 1 Turbidity |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 7 [590 nm] det. 1 unbound Phycocyanin |  | 1.6769 | digits/Âµg/l |  |  | 1.6769
Fingerprint of LED 7 [590 nm] det. 1 Yellow substances |  | 0.22475 | digits/Âµg/l |  |  | 0.22475
Fingerprint of LED 8 [470 nm] det. 1 Green Algae |  | 3.2404 | digits/Âµg/l |  |  | 3.2404
Fingerprint of LED 8 [470 nm] det. 1 Bluegreen |  | 0.30621 | digits/Âµg/l |  |  | 0.30621
Fingerprint of LED 8 [470 nm] det. 1 Diatoms |  | 4.6202 | digits/Âµg/l |  |  | 4.6202
Fingerprint of LED 8 [470 nm] det. 1 Cryptophyta |  | 4.0617 | digits/Âµg/l |  |  | 4.0617
Fingerprint of LED 8 [470 nm] det. 1 Planktothrix |  | 0.8111 | digits/Âµg/l |  |  | 0.8111
Fingerprint of LED 8 [470 nm] det. 1 #6 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 8 [470 nm] det. 1 #7 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 8 [470 nm] det. 1 #8 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 8 [470 nm] det. 1 #9 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 8 [470 nm] det. 1 Turbidity |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 8 [470 nm] det. 1 unbound Phycocyanin |  | 0.028398 | digits/Âµg/l |  |  | 0.028398
Fingerprint of LED 8 [470 nm] det. 1 Yellow substances |  | 9.8385 | digits/Âµg/l |  |  | 9.8385
Fingerprint of LED 1 [700 nm] det. 2 Green Algae |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 1 [700 nm] det. 2 Bluegreen |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 1 [700 nm] det. 2 Diatoms |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 1 [700 nm] det. 2 Cryptophyta |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 1 [700 nm] det. 2 Planktothrix |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 1 [700 nm] det. 2 #6 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 1 [700 nm] det. 2 #7 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 1 [700 nm] det. 2 #8 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 1 [700 nm] det. 2 #9 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 1 [700 nm] det. 2 Turbidity |  | 0.26671 | digits/Âµg/l |  |  | 0.26671
Fingerprint of LED 1 [700 nm] det. 2 unbound Phycocyanin |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 1 [700 nm] det. 2 Yellow substances |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 3 [525 nm] det. 2 Green Algae |  | -0.0048805 | digits/Âµg/l |  |  | -0.0048805
Fingerprint of LED 3 [525 nm] det. 2 Bluegreen |  | 0.21879 | digits/Âµg/l |  |  | 0.21879
Fingerprint of LED 3 [525 nm] det. 2 Diatoms |  | 0.0086236 | digits/Âµg/l |  |  | 0.0086236
Fingerprint of LED 3 [525 nm] det. 2 Cryptophyta |  | 2.3513 | digits/Âµg/l |  |  | 2.3513
Fingerprint of LED 3 [525 nm] det. 2 Planktothrix |  | 1.1733 | digits/Âµg/l |  |  | 1.1733
Fingerprint of LED 3 [525 nm] det. 2 #6 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 3 [525 nm] det. 2 #7 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 3 [525 nm] det. 2 #8 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 3 [525 nm] det. 2 #9 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 3 [525 nm] det. 2 Turbidity |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 3 [525 nm] det. 2 unbound Phycocyanin |  | 0.35338 | digits/Âµg/l |  |  | 0.35338
Fingerprint of LED 3 [525 nm] det. 2 Yellow substances |  | 3.1654 | digits/Âµg/l |  |  | 3.1654
Fingerprint of LED 4 [570 nm] det. 2 Green Algae |  | 0.0025358 | digits/Âµg/l |  |  | 0.0025358
Fingerprint of LED 4 [570 nm] det. 2 Bluegreen |  | 0.2397 | digits/Âµg/l |  |  | 0.2397
Fingerprint of LED 4 [570 nm] det. 2 Diatoms |  | 0.0041241 | digits/Âµg/l |  |  | 0.0041241
Fingerprint of LED 4 [570 nm] det. 2 Cryptophyta |  | 1.4593 | digits/Âµg/l |  |  | 1.4593
Fingerprint of LED 4 [570 nm] det. 2 Planktothrix |  | 0.54995 | digits/Âµg/l |  |  | 0.54995
Fingerprint of LED 4 [570 nm] det. 2 #6 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 4 [570 nm] det. 2 #7 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 4 [570 nm] det. 2 #8 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 4 [570 nm] det. 2 #9 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 4 [570 nm] det. 2 Turbidity |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 4 [570 nm] det. 2 unbound Phycocyanin |  | 0.37198 | digits/Âµg/l |  |  | 0.37198
Fingerprint of LED 4 [570 nm] det. 2 Yellow substances |  | 0.21332 | digits/Âµg/l |  |  | 0.21332
Fingerprint of LED 5 [610 nm] det. 2 Green Algae |  | 0.027878 | digits/Âµg/l |  |  | 0.027878
Fingerprint of LED 5 [610 nm] det. 2 Bluegreen |  | 1.6067 | digits/Âµg/l |  |  | 1.6067
Fingerprint of LED 5 [610 nm] det. 2 Diatoms |  | 0.066813 | digits/Âµg/l |  |  | 0.066813
Fingerprint of LED 5 [610 nm] det. 2 Cryptophyta |  | 2.02 | digits/Âµg/l |  |  | 2.02
Fingerprint of LED 5 [610 nm] det. 2 Planktothrix |  | 0.67318 | digits/Âµg/l |  |  | 0.67318
Fingerprint of LED 5 [610 nm] det. 2 #6 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 5 [610 nm] det. 2 #7 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 5 [610 nm] det. 2 #8 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 5 [610 nm] det. 2 #9 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 5 [610 nm] det. 2 Turbidity |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 5 [610 nm] det. 2 unbound Phycocyanin |  | 2.2011 | digits/Âµg/l |  |  | 2.2011
Fingerprint of LED 5 [610 nm] det. 2 Yellow substances |  | -0.7723 | digits/Âµg/l |  |  | -0.7723
Fingerprint of LED 6 [370 nm] det. 2 Green Algae |  | -0.029187 | digits/Âµg/l |  |  | -0.029187
Fingerprint of LED 6 [370 nm] det. 2 Bluegreen |  | 0.33101 | digits/Âµg/l |  |  | 0.33101
Fingerprint of LED 6 [370 nm] det. 2 Diatoms |  | -0.0091053 | digits/Âµg/l |  |  | -0.0091053
Fingerprint of LED 6 [370 nm] det. 2 Cryptophyta |  | 0.34167 | digits/Âµg/l |  |  | 0.34167
Fingerprint of LED 6 [370 nm] det. 2 Planktothrix |  | 0.2278 | digits/Âµg/l |  |  | 0.2278
Fingerprint of LED 6 [370 nm] det. 2 #6 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 6 [370 nm] det. 2 #7 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 6 [370 nm] det. 2 #8 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 6 [370 nm] det. 2 #9 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 6 [370 nm] det. 2 Turbidity |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 6 [370 nm] det. 2 unbound Phycocyanin |  | 0.40961 | digits/Âµg/l |  |  | 0.40961
Fingerprint of LED 6 [370 nm] det. 2 Yellow substances |  | 25.354 | digits/Âµg/l |  |  | 25.354
Fingerprint of LED 7 [590 nm] det. 2 Green Algae |  | 0.0073756 | digits/Âµg/l |  |  | 0.0073756
Fingerprint of LED 7 [590 nm] det. 2 Bluegreen |  | 1.3214 | digits/Âµg/l |  |  | 1.3214
Fingerprint of LED 7 [590 nm] det. 2 Diatoms |  | 0.014843 | digits/Âµg/l |  |  | 0.014843
Fingerprint of LED 7 [590 nm] det. 2 Cryptophyta |  | 3.468 | digits/Âµg/l |  |  | 3.468
Fingerprint of LED 7 [590 nm] det. 2 Planktothrix |  | 0.87936 | digits/Âµg/l |  |  | 0.87936
Fingerprint of LED 7 [590 nm] det. 2 #6 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 7 [590 nm] det. 2 #7 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 7 [590 nm] det. 2 #8 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 7 [590 nm] det. 2 #9 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 7 [590 nm] det. 2 Turbidity |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 7 [590 nm] det. 2 unbound Phycocyanin |  | 1.9496 | digits/Âµg/l |  |  | 1.9496
Fingerprint of LED 7 [590 nm] det. 2 Yellow substances |  | 1.0055 | digits/Âµg/l |  |  | 1.0055
Fingerprint of LED 8 [470 nm] det. 2 Green Algae |  | -0.014348 | digits/Âµg/l |  |  | -0.014348
Fingerprint of LED 8 [470 nm] det. 2 Bluegreen |  | 0.044611 | digits/Âµg/l |  |  | 0.044611
Fingerprint of LED 8 [470 nm] det. 2 Diatoms |  | 0.0096069 | digits/Âµg/l |  |  | 0.0096069
Fingerprint of LED 8 [470 nm] det. 2 Cryptophyta |  | 0.26029 | digits/Âµg/l |  |  | 0.26029
Fingerprint of LED 8 [470 nm] det. 2 Planktothrix |  | 0.1273 | digits/Âµg/l |  |  | 0.1273
Fingerprint of LED 8 [470 nm] det. 2 #6 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 8 [470 nm] det. 2 #7 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 8 [470 nm] det. 2 #8 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 8 [470 nm] det. 2 #9 |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 8 [470 nm] det. 2 Turbidity |  | 0 | digits/Âµg/l |  |  | 0
Fingerprint of LED 8 [470 nm] det. 2 unbound Phycocyanin |  | 0.027179 | digits/Âµg/l |  |  | 0.027179
Fingerprint of LED 8 [470 nm] det. 2 Yellow substances |  | 10.117 | digits/Âµg/l |  |  | 10.117
Standard deviation of LED 1 [700 nm] det. 1 Green Algae | Standard deviations of the fingerprints of each algae class. This standard deviation refers to variations within one specific algae class. The greater the standard deviation, the less important the measured value for this algae class | 0.1 |  |  |  | 0.1
Standard deviation of LED 1 [700 nm] det. 1 Bluegreen |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 1 [700 nm] det. 1 Diatoms |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 1 [700 nm] det. 1 Cryptophyta |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 1 [700 nm] det. 1 Planktothrix |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 1 [700 nm] det. 1 #6 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 1 [700 nm] det. 1 #7 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 1 [700 nm] det. 1 #8 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 1 [700 nm] det. 1 #9 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 1 [700 nm] det. 1 Turbidity |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 1 [700 nm] det. 1 unbound Phycocyanin |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 1 [700 nm] det. 1 Yellow substances |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 3 [525 nm] det. 1 Green Algae |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 3 [525 nm] det. 1 Bluegreen |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 3 [525 nm] det. 1 Diatoms |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 3 [525 nm] det. 1 Cryptophyta |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 3 [525 nm] det. 1 Planktothrix |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 3 [525 nm] det. 1 #6 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 3 [525 nm] det. 1 #7 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 3 [525 nm] det. 1 #8 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 3 [525 nm] det. 1 #9 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 3 [525 nm] det. 1 Turbidity |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 3 [525 nm] det. 1 unbound Phycocyanin |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 3 [525 nm] det. 1 Yellow substances |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 4 [570 nm] det. 1 Green Algae |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 4 [570 nm] det. 1 Bluegreen |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 4 [570 nm] det. 1 Diatoms |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 4 [570 nm] det. 1 Cryptophyta |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 4 [570 nm] det. 1 Planktothrix |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 4 [570 nm] det. 1 #6 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 4 [570 nm] det. 1 #7 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 4 [570 nm] det. 1 #8 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 4 [570 nm] det. 1 #9 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 4 [570 nm] det. 1 Turbidity |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 4 [570 nm] det. 1 unbound Phycocyanin |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 4 [570 nm] det. 1 Yellow substances |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 5 [610 nm] det. 1 Green Algae |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 5 [610 nm] det. 1 Bluegreen |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 5 [610 nm] det. 1 Diatoms |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 5 [610 nm] det. 1 Cryptophyta |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 5 [610 nm] det. 1 Planktothrix |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 5 [610 nm] det. 1 #6 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 5 [610 nm] det. 1 #7 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 5 [610 nm] det. 1 #8 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 5 [610 nm] det. 1 #9 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 5 [610 nm] det. 1 Turbidity |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 5 [610 nm] det. 1 unbound Phycocyanin |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 5 [610 nm] det. 1 Yellow substances |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 6 [370 nm] det. 1 Green Algae |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 6 [370 nm] det. 1 Bluegreen |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 6 [370 nm] det. 1 Diatoms |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 6 [370 nm] det. 1 Cryptophyta |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 6 [370 nm] det. 1 Planktothrix |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 6 [370 nm] det. 1 #6 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 6 [370 nm] det. 1 #7 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 6 [370 nm] det. 1 #8 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 6 [370 nm] det. 1 #9 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 6 [370 nm] det. 1 Turbidity |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 6 [370 nm] det. 1 unbound Phycocyanin |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 6 [370 nm] det. 1 Yellow substances |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 7 [590 nm] det. 1 Green Algae |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 7 [590 nm] det. 1 Bluegreen |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 7 [590 nm] det. 1 Diatoms |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 7 [590 nm] det. 1 Cryptophyta |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 7 [590 nm] det. 1 Planktothrix |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 7 [590 nm] det. 1 #6 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 7 [590 nm] det. 1 #7 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 7 [590 nm] det. 1 #8 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 7 [590 nm] det. 1 #9 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 7 [590 nm] det. 1 Turbidity |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 7 [590 nm] det. 1 unbound Phycocyanin |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 7 [590 nm] det. 1 Yellow substances |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 8 [470 nm] det. 1 Green Algae |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 8 [470 nm] det. 1 Bluegreen |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 8 [470 nm] det. 1 Diatoms |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 8 [470 nm] det. 1 Cryptophyta |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 8 [470 nm] det. 1 Planktothrix |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 8 [470 nm] det. 1 #6 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 8 [470 nm] det. 1 #7 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 8 [470 nm] det. 1 #8 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 8 [470 nm] det. 1 #9 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 8 [470 nm] det. 1 Turbidity |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 8 [470 nm] det. 1 unbound Phycocyanin |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 8 [470 nm] det. 1 Yellow substances |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 1 [700 nm] det. 2 Green Algae |  | 100 |  |  |  | 100
Standard deviation of LED 1 [700 nm] det. 2 Bluegreen |  | 100 |  |  |  | 100
Standard deviation of LED 1 [700 nm] det. 2 Diatoms |  | 100 |  |  |  | 100
Standard deviation of LED 1 [700 nm] det. 2 Cryptophyta |  | 100 |  |  |  | 100
Standard deviation of LED 1 [700 nm] det. 2 Planktothrix |  | 100 |  |  |  | 100
Standard deviation of LED 1 [700 nm] det. 2 #6 |  | 100 |  |  |  | 100
Standard deviation of LED 1 [700 nm] det. 2 #7 |  | 100 |  |  |  | 100
Standard deviation of LED 1 [700 nm] det. 2 #8 |  | 100 |  |  |  | 100
Standard deviation of LED 1 [700 nm] det. 2 #9 |  | 100 |  |  |  | 100
Standard deviation of LED 1 [700 nm] det. 2 Turbidity |  | 100 |  |  |  | 100
Standard deviation of LED 1 [700 nm] det. 2 unbound Phycocyanin |  | 100 |  |  |  | 100
Standard deviation of LED 1 [700 nm] det. 2 Yellow substances |  | 100 |  |  |  | 100
Standard deviation of LED 3 [525 nm] det. 2 Green Algae |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 3 [525 nm] det. 2 Bluegreen |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 3 [525 nm] det. 2 Diatoms |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 3 [525 nm] det. 2 Cryptophyta |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 3 [525 nm] det. 2 Planktothrix |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 3 [525 nm] det. 2 #6 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 3 [525 nm] det. 2 #7 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 3 [525 nm] det. 2 #8 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 3 [525 nm] det. 2 #9 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 3 [525 nm] det. 2 Turbidity |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 3 [525 nm] det. 2 unbound Phycocyanin |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 3 [525 nm] det. 2 Yellow substances |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 4 [570 nm] det. 2 Green Algae |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 4 [570 nm] det. 2 Bluegreen |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 4 [570 nm] det. 2 Diatoms |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 4 [570 nm] det. 2 Cryptophyta |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 4 [570 nm] det. 2 Planktothrix |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 4 [570 nm] det. 2 #6 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 4 [570 nm] det. 2 #7 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 4 [570 nm] det. 2 #8 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 4 [570 nm] det. 2 #9 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 4 [570 nm] det. 2 Turbidity |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 4 [570 nm] det. 2 unbound Phycocyanin |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 4 [570 nm] det. 2 Yellow substances |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 5 [610 nm] det. 2 Green Algae |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 5 [610 nm] det. 2 Bluegreen |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 5 [610 nm] det. 2 Diatoms |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 5 [610 nm] det. 2 Cryptophyta |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 5 [610 nm] det. 2 Planktothrix |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 5 [610 nm] det. 2 #6 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 5 [610 nm] det. 2 #7 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 5 [610 nm] det. 2 #8 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 5 [610 nm] det. 2 #9 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 5 [610 nm] det. 2 Turbidity |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 5 [610 nm] det. 2 unbound Phycocyanin |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 5 [610 nm] det. 2 Yellow substances |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 6 [370 nm] det. 2 Green Algae |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 6 [370 nm] det. 2 Bluegreen |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 6 [370 nm] det. 2 Diatoms |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 6 [370 nm] det. 2 Cryptophyta |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 6 [370 nm] det. 2 Planktothrix |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 6 [370 nm] det. 2 #6 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 6 [370 nm] det. 2 #7 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 6 [370 nm] det. 2 #8 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 6 [370 nm] det. 2 #9 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 6 [370 nm] det. 2 Turbidity |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 6 [370 nm] det. 2 unbound Phycocyanin |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 6 [370 nm] det. 2 Yellow substances |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 7 [590 nm] det. 2 Green Algae |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 7 [590 nm] det. 2 Bluegreen |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 7 [590 nm] det. 2 Diatoms |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 7 [590 nm] det. 2 Cryptophyta |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 7 [590 nm] det. 2 Planktothrix |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 7 [590 nm] det. 2 #6 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 7 [590 nm] det. 2 #7 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 7 [590 nm] det. 2 #8 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 7 [590 nm] det. 2 #9 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 7 [590 nm] det. 2 Turbidity |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 7 [590 nm] det. 2 unbound Phycocyanin |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 7 [590 nm] det. 2 Yellow substances |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 8 [470 nm] det. 2 Green Algae |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 8 [470 nm] det. 2 Bluegreen |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 8 [470 nm] det. 2 Diatoms |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 8 [470 nm] det. 2 Cryptophyta |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 8 [470 nm] det. 2 Planktothrix |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 8 [470 nm] det. 2 #6 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 8 [470 nm] det. 2 #7 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 8 [470 nm] det. 2 #8 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 8 [470 nm] det. 2 #9 |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 8 [470 nm] det. 2 Turbidity |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 8 [470 nm] det. 2 unbound Phycocyanin |  | 0.1 |  |  |  | 0.1
Standard deviation of LED 8 [470 nm] det. 2 Yellow substances |  | 0.1 |  |  |  | 0.1
Date of calibration | Date of the last calibration at bbe Moldaenke | 20200409162511 |  |  |  | 20200409162511
Global correction factor | Set to 1 by default. All algae concentrations are multiplied by this factor | 1 |  |  |  | 1
Cell factor Green Algae | Factor to convert the concentration of the algae classes to cell counts | 530000 |  |  |  | 530000
Cell factor Bluegreen |  | 1000000 |  |  |  | 1000000
Cell factor Diatoms |  | 450000 |  |  |  | 450000
Cell factor Cryptophyta |  | 30000 |  |  |  | 30000
Cell factor Planktothrix |  | 1000000 |  |  |  | 1000000
Cell factor #6 |  | 1000000 |  |  |  | 1000000
Cell factor #7 |  | 1000000 |  |  |  | 1000000
Cell factor #8 |  | 1000000 |  |  |  | 1000000
Cell factor #9
Cell factor Turbidity
Cell factor unbound Phycocyanin
Cell factor Yellow substances
Unit parameters | Here the algae class specific unit could be changed if differing from standard μg/L
Name | Explanation | Value | Unit | Questions | Notes | Default Value
Green Algae concentration
Bluegreen concentration
Diatoms concentration
Cryptophyta concentration
Planktothrix concentration
#6 concentration
#7 concentration
#8 concentration |  | r.u. |  |  |  | r.u.
#9 concentration
Turbidity concentration |  | FTU |  |  |  | FTU
unbound Phycocyanin concentration
Yellow substances concentration |  | r.u. |  |  |  | r.u.
Green Algae cellcounts
Bluegreen cellcounts
Diatoms cellcounts
Cryptophyta cellcounts
Planktothrix cellcounts
#6 cellcounts
#7 cellcounts
#8 cellcounts |  | r.u. |  |  |  | r.u.
#9 cellcounts
Turbidity cellcounts |  | FTU |  |  |  | FTU
unbound Phycocyanin cellcounts
Yellow substances cellcounts |  | r.u. |  |  |  | r.u.
Measurement parameters
Name | Explanation | Value | Unit | Questions | Notes | Default Value
Air pressure | Displays the air pressure. This parameter is set when clicking the “Set air pressure“ button before starting the measurement. The parameter is used to calculate the depth. | 1030 | mbar |  |  | 1000
DA value of LED 1 [700 nm] | Current settings of the LED controller | 2640 | digits |  |  | 2640
DA value of LED 3 [525 nm] |  | 2409 | digits |  |  | 2451
DA value of LED 4 [570 nm] |  | 2235 | digits |  |  | 2382
DA value of LED 5 [610 nm] |  | 2495 | digits |  |  | 2664
DA value of LED 6 [370 nm] |  | 2572 | digits |  |  | 2571
DA value of LED 7 [590 nm] |  | 2385 | digits |  |  | 2550
DA value of LED 8 [470 nm] |  | 2194 | digits |  |  | 2296
Pressure offset | Displays the results of the calibrations of the measurands | 0 | digits |  |  | 0
Pressure gradient |  | 0.00835 | bar/digit |  |  | 0.00835
Salinity | Enter the salinity of the water to enhance the calculation of the depth | 0 | psu |  |  | 0
Battery offset |  | 0 | digits |  |  | 0
Battery gradient |  | 0.015444 | V/digit |  |  | 0.015444
Sample temperature offset |  | 319 | digits |  |  | 319
Sample temperature gradient |  | 0.061998 | Â°C/digit |  |  | 0.061998
Required value of LED 1 [700 nm] | Values for the basic calibration of the PhycoProbe | 0 | digits |  |  | 0
Required value of LED 3 [525 nm] |  | -479.7 | digits |  |  | -479.7
Required value of LED 4 [570 nm] |  | -97.63 | digits |  |  | -97.63
Required value of LED 5 [610 nm] |  | -883.05 | digits |  |  | -883.05
Required value of LED 6 [370 nm] |  | -187.65 | digits |  |  | -187.65
Required value of LED 7 [590 nm] |  | -355.41 | digits |  |  | -355.41
Required value of LED 8 [470 nm] |  | -384.92 | digits |  |  | -384.92
Temp. of LED offset |  | 0 | digits |  |  | 0
Temp. of LED gradient |  | 0.037905 | Â°C/digit |  |  | 0.037905
Temperature correction of LED 0 to 5 Â°C | Temperature correction factors for the given temperature ranges | 1 |  |  |  | 1
Temperature correction of LED 5 to 10 Â°C |  | 1 |  |  |  | 1
Temperature correction of LED 10 to 15 Â°C |  | 1 |  |  |  | 1
Temperature correction of LED 15 to 20 Â°C |  | 1 |  |  |  | 1
Temperature correction of LED 20 to 25 Â°C |  | 1 |  |  |  | 1
Temperature correction of LED 25 to 30 Â°C |  | 1 |  |  |  | 1
Temperature correction of LED 30 to 35 Â°C |  | 1 |  |  |  | 1
Temperature correction of LED 35 to 40 Â°C |  | 1 |  |  |  | 1
Temperature correction of LED 40 to 45 Â°C |  | 1 |  |  |  | 1
Temperature correction of LED 45 to 50 Â°C |  | 1 |  |  |  | 1
Temperature correction of LED > 50 Â°C |  | 1 |  |  |  | 1
Temp. of sensor offset |  | 0 | digits |  |  | 0
Temp. of sensor gradient |  | 0.13094 | Â°C/digit |  |  | 0.13094
Temperature correction of sensor 0 to 5 Â°C |  | 0.964 |  |  |  | 0.964
Temperature correction of sensor 5 to 10 Â°C |  | 0.973 |  |  |  | 0.973
Temperature correction of sensor 10 to 15 Â°C |  | 0.982 |  |  |  | 0.982
Temperature correction of sensor 15 to 20 Â°C |  | 0.991 |  |  |  | 0.991
Temperature correction of sensor 20 to 25 Â°C |  | 1 |  |  |  | 1
Temperature correction of sensor 25 to 30 Â°C |  | 1.02 |  |  |  | 1.02
Temperature correction of sensor 30 to 35 Â°C |  | 1.041 |  |  |  | 1.041
Temperature correction of sensor 35 to 40 Â°C |  | 1.062 |  |  |  | 1.062
Temperature correction of sensor 40 to 45 Â°C |  | 1.083 |  |  |  | 1.083
Temperature correction of sensor 45 to 50 Â°C |  | 1.104 |  |  |  | 1.104
Temperature correction of sensor > 50 Â°C |  | 1.125 |  |  |  | 1.125
Turbidity parameters
Name | Explanation | Value | Unit | Questions | Notes | Default Value
Turbidity compensation | Setting to activate and deactivate the turbidity compensation | on |  |  |  | on
FTU reflection compensation LED 1 [700 nm] det. 1 | Factors to adjust the turbidity compensation that consider light scattering effects as well as inner filtering effects | 0 |  |  |  | 0
FTU reflection compensation LED 3 [525 nm] det. 1 |  | 0.0066 |  |  |  | 0.0066
FTU reflection compensation LED 4 [570 nm] det. 1 |  | 1e-10 |  |  |  | 1e-10
FTU reflection compensation LED 5 [610 nm] det. 1 |  | 1e-10 |  |  |  | 1e-10
FTU reflection compensation LED 6 [370 nm] det. 1 |  | 0.0369 |  |  |  | 0.0369
FTU reflection compensation LED 7 [590 nm] det. 1 |  | 0.0031 |  |  |  | 0.0031
FTU reflection compensation LED 8 [470 nm] det. 1 |  | 0.0671 |  |  |  | 0.0671
FTU mitigation compensation LED 1 [700 nm] det. 1 |  | 0 |  |  |  | 0
FTU mitigation compensation LED 3 [525 nm] det. 1 |  | 0.051 |  |  |  | 0.051
FTU mitigation compensation LED 4 [570 nm] det. 1 |  | 0.048 |  |  |  | 0.048
FTU mitigation compensation LED 5 [610 nm] det. 1 |  | 0.03 |  |  |  | 0.03
FTU mitigation compensation LED 6 [370 nm] det. 1 |  | 0.323 |  |  |  | 0.323
FTU mitigation compensation LED 7 [590 nm] det. 1 |  | 0.036 |  |  |  | 0.036
FTU mitigation compensation LED 8 [470 nm] det. 1 |  | 0.083 |  |  |  | 0.083
FTU reflection compensation LED 1 [700 nm] det. 2 |  | 0 |  |  |  | 0
FTU reflection compensation LED 3 [525 nm] det. 2 |  | 0.0311 |  |  |  | 0.0311
FTU reflection compensation LED 4 [570 nm] det. 2 |  | 1e-10 |  |  |  | 1e-10
FTU reflection compensation LED 5 [610 nm] det. 2 |  | 0.0707 |  |  |  | 0.0707
FTU reflection compensation LED 6 [370 nm] det. 2 |  | 0.0198 |  |  |  | 0.0198
FTU reflection compensation LED 7 [590 nm] det. 2 |  | 0.0243 |  |  |  | 0.0243
FTU reflection compensation LED 8 [470 nm] det. 2 |  | 0.0428 |  |  |  | 0.0428
FTU mitigation compensation LED 1 [700 nm] det. 2 |  | 0 |  |  |  | 0
FTU mitigation compensation LED 3 [525 nm] det. 2 |  | 0.169 |  |  |  | 0.169
FTU mitigation compensation LED 4 [570 nm] det. 2 |  | 0.148 |  |  |  | 0.148
FTU mitigation compensation LED 5 [610 nm] det. 2 |  | 0.151 |  |  |  | 0.151
FTU mitigation compensation LED 6 [370 nm] det. 2 |  | 0.493 |  |  |  | 0.493
FTU mitigation compensation LED 7 [590 nm] det. 2 |  | 0.117 |  |  |  | 0.117
FTU mitigation compensation LED 8 [470 nm] det. 2 |  | 0.271 |  |  |  | 0.271